Data quality control and tools in passive seismic experiments exemplified on Czech broad-band seismic pool MOBNET in the AlpArray collaborative project

Luděk Vecsey1, Jaroslava Plomerová1, Petr Jedlička1, Helena Munzarová1, Vladislav Babuška1 and the AlpArray Working Group*

1Institute of Geophysics of the Czech Academy of Sciences, 14131 Prague, Czech Republic

*www.alparray.ethz.ch

Correspondence to: Luděk Vecsey (vecsey@ig.cas.cz)

Abstract. This paper focuses on major issues related to data reliability and network performance of twenty broad-band stations of the Czech MOBNET pool within the AlpArray seismic experiments. Currently used high-resolution seismological applications require high-quality data recorded for a sufficiently long time interval at seismological observatories and during the whole time of operation of temporary stations. In this paper we present new hardware and software tools, we have been developing during the last two decades while analysing data from several international passive experiments. The new tools help to assure the high-quality standard of broad-band seismic data and eliminate potential errors before supplying data to seismological centres. Special attention is paid to crucial issues like the detection of sensor mis-orientation, timing problems, interchange of record components and/or their polarity reversal, sensor mass centring, or anomalous channel amplitudes due to, e.g., imperfect gain.  Thorough data-quality control should represent an integral constituent of seismic data recording, pre-processing and archiving, especially for data from temporary stations in passive seismic experiments. Large international seismic experiments require enormous efforts of scientists from different countries and institutions to gather hundreds of stations to be deployed in the field during a limited time period. In this paper, we demonstrate the beneficial effects of the procedures we have developed for having reliable large set of high-quality data from each group participating in field experiments. The presented tools can be applied manually or automatically on data from any seismic network.

# 1 Introduction

The long-term experience of some of the authors with passive-experiment data processing encourage the team to summarize tools we have been using for testing and improving seismic data quality and which might be of interest for a broader community. Necessity of data quality control is evident nowadays and several procedures are applied automatically in data centres, e.g. software MUSTANG in IRIS, which identify data errors in the centre databases. Our endeavour is to identify data errors and correct them when possible, before supplying data to the centres. Data quality control before experimental data archiving is of great importance.

Data from passive seismic experiments of different lateral extend and a dense station distribution, became crucial source of information for the modern Earth interior researches. The USArray (www.usarray.org) or IberArray (iberarray.ictja.csic.es; Díaz et al., 2010) represent the large-scale temporary networks, whereas, e.g., TRANSALP (Lippitsch et al., 2003) or BOHEMA (Plomerová et al., 2007) belong to small-size passive experiments in central Europe. Participants of the AlpArray project, the European collaborative geoscience initiative, deployed the largest network of temporary broad-band station ever performed in Europe (AlpArray Seismic Network, 2015; www.alparray.ethz.ch). The project focuses on the structure and evolution of the lithosphere-asthenosphere system beneath the greater Alpine area – the Alps and their forelands and thus it requires an intensive international cooperation in data gathering. The northern foreland of the Alps is formed by the Bohemian Massif, the easternmost outcrop of the Variscan belt of the European plate. The project makes use of seismological as well as associated Earth science data for better understanding the geodynamics of the greater Alpine area and its seismic hazard. The area, studied by generations of geoscientists, comprises the orogenic system, where two large plates (Europe and Africa) have converged and interacted over time with several micro-plates of oceanic and continental provenances (Kissling et al., 2006; Handy et al., 2010 for reviews). Besides the Alpine structure itself, the Alps-Apennines, Alps-Dinarides and Alps-Bohemian Massif contacts in depth are of the particular interest within the AlpArray study. In addition to structural studies related to the orogenic system dominating Europe with the use of associated Earth-science data (such as gravity, electro-magnetics, geology, etc.), several other topics as seismotectonics and earthquake hazard belong to the core of the project. Various seismological methods, including tomography, ambient noise analysis and receiver functions, considering anisotropy in all three types of investigations, as well as in shear-wave splitting analyses, will be applied. The depth range of scientific investigations encompasses the crust and the mantle lithosphere, down to the lithosphere-asthenosphere boundary (LAB), as well as the sub-lithospheric upper mantle.

To achieve objectives of the project, it is necessary to apply various geological/geophysical imaging methods on data recorded by a homogeneous network of broad-band (BB) seismic stations in the greater Alpine area (Fig. 1). Though the area is in some parts densely covered by permanent seismic observatories, their distribution is far from being homogeneous. Therefore, the distribution of ~360 existing permanent stations have been complemented by ~260 temporary BB stations to create a relatively dense network of unprecedented large scale in Europe, with homogeneous station spacing of about 52 km. The station spacing and station location is designed in such a way that for any site in the Alpine region there is always a station of the AlpArray Seismic Network (AASN) at a distance up to ~30 km. The temporary seismic network of such large extent requires intensive collaboration between many institutions (currently more than 45 institutions from 17 countries), combination of individual national/institutional seismic pools of temporary stations and coordination of their deployment, keeping the high-level maintenance and experienced handling. Thanks to the large extent of the array and density of the stations, results from seismic tomography and several other techniques applied on data collected during the unique passive experiment will shed light on the detailed 3D architecture of the crust and upper mantle from the Earth’s surface down to ~600 km of this extremely complicated orogenic region.

The AlpArray area, set as a region delimited by a 250 km distance from the 800 m altitude isoline surrounding the Alps, covers a large portion of the Czech (CZ) part of the Bohemian Massif (BM). Ten BB observatories of the Czech Regional Seismological Network (CRSN), one permanent BB station of the West Bohemian Network (WEBNET) along with 20 temporary BB stations from pool of seismic stations MOBNET (MOBile NETwork) of the Institute of Geophysics, Czech Academy of Sciences (IG CAS), cover the area with the spacing required (Fig. 1). Apart of the AASN, the backbone of the AlpArray project, the MOBNET stations were involved in the complementary project AlpArray-EASI (coded XT in the European Integrated Data Archive EIDA). Additional 10 stations operate since June 2017 in another complementary project AlpArray-IVREA (coded XK in the European Integrated Data Archive EIDA). The Czech team is responsible for the deployment and maintenance of the MOBNET stations included in the Czech part of the AASN (coded Z3 in the European Integrated Data Archive (EIDA) system, www.fdsn.org/networks/detail/Z3_2015/), as well as for completeness and correctness of seismic data recorded in the complementary experiment AlpArray-EASI and transferred to the EIDA centres.

The main purpose of this paper is to describe technical parameters of the MOBNET stations, to present newly developed control units for setting sensor and data acquisition systems (DAS) and to document significance of careful data-quality control, which could help other groups not only those involved in the AlpArray project in preparing their seismic data for archiving. Special attention is paid to detection of sensor mis-orientation, timing problems, interchanged of components and/or their polarity reversal, mass centring problems, or anomalous channel amplitudes due to, e.g., an imperfect gain. Keeping the high quality of archived seismological data is crucial for success of the AlpArray project, as well as of any passive seismic experiment.

# 2 Deployment of MOBNET stations within the AlpArray project

Twenty stations of the MOBNET pool have been deployed in the Bohemian Massif (BM) since August 2015, as a part of the AASN (Fig. 2). The first integration of the twenty MOBNET stations into the AlpArray project had been realized during the Eastern Alpine Seismic Investigation (EASI) project, the first implemented AlpArray Complementary Experiment (Table 1, see Fig. 1). The  AlpArray-EASI transect was composed of 55 broadband seismic stations. They operated from July 2014 to October 2015 and were configured in a zig-zag pattern on either side of the central longitude line of 13.35° E, with the north-south distance between stations of 10 km. The transect spanned a 540 km long region, between the Erzgebirge Mts. at the Czech-German border in the North and the Adriatic Sea, near Trieste in the South. The distance of each station to either side of the central line was ~6 km. We followed the general recommendations of the Technical strategy of the AlpArray project (www.alparray.ethz.ch), including a temperature insulation of sensors. We kept the stations within 1.5 km of the target location, if topographic, field and infrastructure conditions allowed.

The northernmost stations AAE01-AAE20 (Fig. 2) of the MOBNET pool involved in the AlpArray–EASI network were equipped mostly with the STS-2 seismometers, two CMG-3T and three CMG-3ESP seismometers, and the GAIA DAS. The stations were installed preferably in vaults of castles/chateau, churches, or suitable abandoned buildings. Figure 3 shows an example of a station location, seismometer installation, quality of the site, noise level, etc. Supplements S1-S19 give the same detailed information for remaining 19 MOBNET stations employed in the AlpArray-EASI network. Keeping notation of Molinari et al. (2016), we can characterize the locations as of urban free-field site, only exceptionally as of building site (Table 1). The stations ran at the autonomous regime and reported daily their state-of-health (SOH) in the SMS messages. Altogether, we recorded 280 GB of data stored in the  Miniseed format which contribute to the AlpArray-EASI studies.  including tomography, ambient noise analysis and receiver functions, considering anisotropy in all three types of investigations, as well as in shear-wave splitting analyses. The depth range of scientific investigations encompasses the crust and the mantle lithosphere, down to the lithosphere-asthenosphere boundary (LAB).

After the end of the AlpArray-EASI field measurement in August 2015, twenty MOBNET stations were re-installed in the Bohemian Massif (BM) as a part of the newly designed AASN (see Fig. 2). With the exception of A090A, the stations operate offline. Data from the offline stations are recorded on flash cards with the capacity exceeding at least 4 times space needed for data sampled at rate of 100 sps and collected in three-month intervals to be checked and supplied to the ODC-EIDA node. Similarly to the AlpArray-EASI transect, most of the AASN-CZ sites are classified as urban free-field types (Table 1). Figure 4 shows installation of one of the MOBNET stations and Supplements S20-S38 give a detailed information for remaining 19 MOBNET stations employed in the AASN of the AlpArray project. Though the region of the BM is densely populated with local industrial and agricultural sources of high-frequency noise, most of the stations meet requested noise limits (Peterson, 1993) as it is shown in Figure 4 for station A076A (see also S20-S38, and Fig. 8). Noise exceeds the limit on vertical components at long-period range (T > 100 s, e.g., S33) only at about 30 % of stations. Some of the stations exhibit distinct seasonal variations of noise level, which results in exceeding the noise limit in the long-period range on horizontal components (Fig. 5), (Wolin et al., 2015).

Data from the Czech temporary stations, with the access restricted according to the AlpArray rules, are transferred to ODC (www.orfeus-eu.org/data/eida/nodes/), while data from the Czech permanent stations with open access are continuously being stored in the GEOFON (geofon.gfz-potsdam.de).  Figure 6 shows current status of data availability from the MOBNET stations included in the AlpArray passive field experiments. In case of the AlpArray-EASI complementary project, ran for 15 months in 2014-2015, we retrieved 96 % of the data at each station, on average (Fig. 6a). As concerns the ongoing AlpArray project, the data completeness is 99 % for the MOBNET stations included in the AASN, whose data have been downloaded by September 2017. Several gaps in data were caused by summer thunderstorms that damaged electrical supplies (Fig. 6b). Though almost all our stations operate offline, the data completeness for MOBNET stations in the AASN is similar to that for stations of the Austrian or Swiss parts of the AASN with an online data transmission (Fuchs et al., 2016; Molinari et al., 2016).

# 3 Seismometer and GAIA control and calibration devices

Our broad-band temporary stations involved in the AlpArray project are equipped mostly with broad-band seismometers STS-2 and several GURALP CMG (Table 1), and with data acquisition systems GAIA developed by the VISTEC company (www.vistec.cz).  Hardware of the GAIA data acquisition systems supports seismometer control, but the firmware we use does not allow it. First, our AlpArray stations operate autonomously (they are not connected to Internet) and second, a technician servicing a station often does not need to use a computer for data collection. For these reasons and to assure a high-degree reliability of the seismometer-DAS pairs performance, we have developed four special control devices for seismometers of different types and one for the GAIA DAS. In general, these boxes generate pulses into the systems and compare amplitudes of the input and output signals. The devices enable to calibrate sensors and data acquisition systems, as well as to check in-situ gain of all individual components and polarity of the recorded signal. The hardware check facilitates identification/verification of any malfunction of the systems and enables their immediate treatment, often directly in the field. These devices have been developed in response to our experience achieved during preceding experiments. We applied the devices during station installations, regular servicing, or in case of station malfunction detected by a software quality check. In the future, they can be used before station deployment together with the Huddle test of the instruments. Nevertheless, some malfunctioning can occur during station operation and therefore, regular checks during station services are recommended.

## 3.1 Guralp host box (CMG-3T and CMG-3ESP(C))

Guralp host box developed in our laboratory (Fig. 7a) becomes an integral constituent of the CMG-3T and CMG-3ESP(C) seismometers. It connects the seismometer and the GAIA DAS and it is an analogy of the standard hand-held unit of the Guralp company provenience, or the host box of the STS-2 seismometer. The standard Guralp host box allows fundamental handlings of the seismometer, namely lock/unlock function of pendulums, their centring, and a calibration with the use of an external signal. On top of the fundamental handlings, the Guralp host box developed in our laboratory enables to apply ,the built-in source of calibration signals (Dirac and rectangular pulse functions). Busy LED light informs about a state of the seismometer. The host box is equipped with a connector for the Guralp control and calibration unit (see Sect. 3.2), or for a remote seismometer control (e.g., via GSM).

## 3.2 Guralp control and calibration unit (CMG-3T and CMG-3ESP(C))

This device (Fig. 7b) enables to display positions of the pendulums and to calibrate a seismometer by the unit rectangular pulse signal or the Dirac delta pulse. It has also an input for external calibrating signal of an arbitrary shape. Polarity of the calibrating signal can be changed and the signal size can be altered in two levels. There is a rotary switch between the calibration mode and the display mode of pendulum positions of the Z, NS and EW components. A push button centres the pendulums. Guralp control and calibration unit is plugged into the Guralp host box connector.

## 3.3 Guralp centring unit (CMG-40T)

Guralp centring unit (Fig. 7c) was developed for seismometer pendulums without electronic centring, e.g., CMG-40T. The unit displays pendulum positions of individual components and thus enables their manual centring. For the pendulum position checking, it is necessary to disconnect the seismometer from the DAS and to connect the Guralp centring unit. Deviation of the pendulum from the central position is proportional to the mass position voltage. Position of the pendulums of the Z, NS, EW components is selected by a switch. Pendulum centring requires the mass position voltage close to zero. The unit has a built-in accumulator, which supplies energy to seismometer during the control. The accumulator voltage is measured in the fourth position of the switch. In case of insufficient accumulator capacity, the accumulator can be plugged-in via an external charger. The Guralp centring unit, developed for seismometers with only a manual pendulum centring, can be used also for pendulum position check of seismometers with electronic control, but in such case, the centring unit does not enable to correct a pendulum possition.

## 3.4 STS-2 control and calibration unit

The STS-2 control and calibration unit (Fig. 7d) has been developed for centring pendulums and for calibrating the Streckeisen STS-2 seismometers.  The device is being connected to the host box provided by the seismometer producer. The host box forms the integral part of the system, through which the STS-2 seismometers are controlled and powered.  The host box has two connectors. The first one is used to connect the digitizer, the second one (marked as "Monitor") serves for a remote control and monitoring of the seismometer via the STS-2 control and calibration unit. This unit displays positions of the pendulums for the U, V and W components, or, it can be switched to show offsets of the standard Z, NS, EW components of the output signal. The unit is equipped with a button of automatic centring of pendulum position (auto-zero push button) connected in parallel to a similar button of the host box. The 120 s /1 s switch of the control and calibration unit changes modes between the broad-band and short-period regimes.

Each of the U, V and W components can be calibrated separately with the unit rectangular signal or the Dirac delta pulse. There is also a switch for an external calibrating signal of an arbitrary shape, e.g., of a sinusoidal signal. If the components are calibrated together, calibration currents and their polarities are chosen so that the output signals (components Z, NS, EW) have the same amplitudes and polarities. This procedure guarantees correct functioning of the seismometer.

## 3.5 GAIA gain and calibration unit

GAIA gain and calibration unit (Fig. 7e) checks and calibrates inputs into the GAIA DAS, but it can be used for calibration of any type of digitizers as well (Kinemetrics, Nanometrics, Reftek, Guralp etc.), after being equipped with corresponding connector reductions. The unit enables to calibrate analogue inputs, to check the correct order of the channels, to determine intensity of cross-talks between the channels, to measure channel amplification and sensitivity (a voltage corresponding to the LSB - least significant bit). Number of channels undergoing calibration and channel polarity can be changed. The calibration is done by a defined voltage jump. For calibration of the analog inputs, we can use a differential, or a single-ended mode. . In the differential mode, the voltage is connected between inputs marked as +IN and -IN. In the single-ended mode, the calibrating voltage is connected to Ground (GND) and +IN or -IN. Built-in generator of saw-shape calibrating voltage serves for an assessment of linearity of the analog-to-digital conversion.

# 4 Data-quality control and assurance

The high level of data quality has to be stable during a long-time interval for seismological observatories and for a whole time of operation of temporary stations within passive experiments. Data-quality control represents the necessary step to achieve and maintain the high-quality level of recorded data. We differentiate (1) in-situ controls with technical equipment, applied during station installation and servicing, and (2) subsequent software controls, applied to downloaded data.

## 4.1 Seismic noise

The measure of seismic ambient noise level is nowadays a standard procedure when searching and selecting sites suitable for station installation. Therefore, we have measured noise at each site before a station installation for a short time. Once a station is installed, the noise level has to be  frequently checked to monitor potential changes in conditions of recordings, or to detect technical problems of a station. According to the AlpArray working group requirements, an average noise level should be 20 dB lower than the New High Noise Model (NHNM; Peterson, 1993) on all components within the 1-10 Hz frequency range. In the long-period range (30-200 s), the same noise level is required only for the vertical component. Because ambient noise is usually higher on the horizontal components, an average noise level is recommended to be only 10 dB less than the NHNM. To follow the ambient noise level, we use the seismic probabilistic power spectral density  procedure (PPSD) by McNamara and Buland (2004) and Custodio et al. (2014), which is a part of the ObsPy module (Krischer et al., 2015).

Figure 8 shows the PPSD medians for all MOBNET stations deployed in the AlpArray-EASI and AASN networks.  While the noise level for periods below 1 s fulfils the noise requirements for most of the stations and for the three components, noise on the horizontal components for periods higher than 10 s is often larger, especially in winter time, but still acceptable for temporary deployments. One has to bear in mind that a compromise between optimal site conditions and the required array geometry has to be accomplished. Thus, at the short period range, we have to accept higher noise level at some sites, where civilization activities are higher (e.g., AAE03 located in an administrative building in a village). Microseisms dominate a period interval of 1-10s in central Europe and also increase in winter. The broad-band seismometers are sensitive to several external effts, especially in the range of longer periods. The most significant of them are temperature changes, either diurnal or seasonal, as well as pressure fluctuations. An enhanced insulation of seismometers might decrease the effects, particularly on the horizontal components. Therefore, seismometer insulation plays an important role to ensure the high-quality data. Majority of our stations are installed in vaults with only small temperature variations, which could directly influence the seismometer pendulums. On the other hand, there are also indirect effects of temperature, particularly an inclination of bedrocks or buildings. Most of our stations are equipped with the STS2 seismometers with three pendulums oriented in 120° each to other. There is no reason, why two reconstructed horizontal components should be affected by direct temperature changes more than the Z component reconstructed from movements of the three pendulums. Only two of our stations, A090A and A078A, are located outside a building, where temperature changes can be significant. These stations exhibit the largest noise at the long-period range, along with the A082A located above loess around the Elbe river. However, we also observe similarly large long-period noise at station A081A, which is in a vault and well insulated. The lowest noise at longer periods is observed at station A076A, where the seismometer is in about 3.5m deep shaft dug in a rock below the building and thus with minimal temperature variations or other jamming. On the other hand, the seismometer of station A084A is located also in almost ideal conditions inside a castle, in a vault space near a rock outcrop, but in spite of that the station exhibits relatively large noise in the long-period range. Noise is generally larger on the EW components than on the NS components in our region (Fig. 8). Sources of these effects are complex and not all of them are clear. Possibly dominated western winds can increase noise on the EW components. Though we are not able to determine all sources of long-period noise, we can exploit the difference in noise levels on horizontal and vertical components as one of tools to decipher potential interchange of the components, as we describe below.

## 4.2 Sensor orientation

Exact orientation of seismometers in the geographic co-ordinate system is one of the most important tasks during station installations. Mis-oriented sensors affect negatively results of procedures based on modern three component seismological observations and can lead to false interpretations (Ekström and Busby, 2008; Vecsey et al., 2014; Wang et al., 2016). Determination of the northward direction has been routinely done for years with the use of standard compass, with the best accuracy of ±5° in case of no magnetic disturbances in nearby surroundings. However, such accuracy is no more sufficient. The top-level current practice is to orient seismometers with the use of the high-precision optical gyrocompass measurement during a station installation and to repeat the measurements during station services. Repeated measurements are desirable to avoid any seismometer mis-orientation resulting from, e.g., an accidental shift of sensors by a person or an animal, as well as due to a nearby strong lightning, which all we have experienced. The higher accuracy in orientation of seismometers towards the North can be achieved only with the optical gyrocompass. But the device is expensive and therefore, still not very commonly used. Being aware of that, we have provided the gyrocompass for measurements in other regional subarrays (e.g., in Slovakia, Austria, Hungary) of the AlpArray project to achieve the correct sensor orientation.

To determine correct sensor orientations, one can use the Rayleigh-wave polarization-angle method (e.g. Stachnik et al., 2012), in which differences between the Rayleigh-wave polarizations and their theoretical back-azimuths are plotted in dependence on event origin times. Of course, this method cannot be as precise as measurements with the optical gyrocompass. Rueda and Mezcua (2015) found only 1-5º difference between the North estimate by the Rayleigh-wave polarization method and the gyrocompass measurements of the North for long-term data series at observatories. In case of shorter time intervals the accuracy of the Rayleigh-wave polarization method is low (can exceed 10º) and thus only large sensor mis-orientation can be detected. For determining an exact moment of the change of sensor orientation, the Rayleigh-wave polarization-angle method can be combined with plots of daily amplitude-means. After determining a day, when the sensor happens to be accidentally mis-oriented, one has to search sudden amplitude changes in the data, which cannot be connected with seismic signal.

When installing our stations for the AlpArray-EASI transect, we oriented the seismometers carefully, but only with the use of a standard compass. Later we checked the orientation of all sensors with a optic gyrocompass. We have found deviations larger than 5° from the true North at 9 of 20 stations (Table 2, 1st measurements) and extremely large deviations in orientation at two of them (AAE13 N=282° and AAE04 N=341°). The first deviation of -78º in the North determination can be attributed to an error which occurs when the STS2 sensor orienting rod is installed in direction towards the North instead of towards the East. Reason of the wrong seismometer orientation in the AAE04 station is unknown. Seismometers at two other stations, AAE13 and AAE18, changed significantly their orientation during the experiment by 8° and 7° (the 5th column in Table 2), respectively. We have used the Rayleigh-wave polarization-angle method for a rough estimate of a moment, when orientation of the sensors has been changed, and daily means and signal plots for setting exact time of the sensor re-orientations.

Sensors of all our stations involved in the currently running AASN network (A071-A090) have been installed with the use of our gyrocompass and their orientation is regularly checked. During about a one-year period of the array operation, we have already recorded three unwanted changes of sensor orientation due to a human intervention. Besides the necessary sensor re-orientation on spot, previous inaccuracies in sensor orientations have been corrected in the metadata. In case that a deviation in the seismometer orientation is larger than 5º, the horizontal N and E components are renamed to components 2 and 3, respectively, according to the SEED Reference Manual (FDSN, 2012, http://www.fdsn.org/media/_s/publications/SEEDManual_V2.4.pdf).

## 4.3 Timing issues

Correct timing is crucial for studies based on exact arrival times of seismic waves. Incorrect time decreases accuracy of picking individual phases, causes false phase identification or a loss of data at all. Timing errors of 1 s or smaller are not clearly evident during routine seismological analyses, but can be revealed from station “log” files, if provided by the registration system, from carefully kept service sheets, or from headers of data. Existing time gaps and overlaps can be calculated from the time of the first sample, number of samples and sampling rate in each  block.

Here we address three important timing problems: the leap second recorded with a delay, switch between the UTC and GPS times and  malfunction of an oscillator tuning the station time or loss of time synchronization for a time period.

The first item – the leap second is introduced into the Coordinated Universal Time (UTC) usually once or twice per year in order to keep the UTC day time close to the mean solar time. The leap second is usually applied at midnight, while clocks in data acquisition systems are being synchronized later, e.g., with a 30-90 minute delay. Moreover, the leap-second correction is applied at individual stations differently, because the times of their synchronizations differ. It is thus necessary to shift the leap second right to the midnight for all temporary stations before data archiving. Surprisingly, we have found a case, when even a permanent observatory, out of Bohemian Massif, kept the uncorrected time for about one month.

The second item – the switch between the UTC and GPS times, can arise due to a wrong synchronization of the inner time (UTC) of a station and the GPS time. This can happen when the coordinated universal time in the “almanac”, transmitted by satellites, disappears from the memory of a station for some reasons (e.g., low voltage of inner battery, incorrect satellite signal recorded, etc.). Existing time gaps and overlaps in Miniseed data can be calculated from the time of the first sample, number of samples and sampling rate in each Miniseed block. Then appropriate time shift is applied in Miniseed data for the identified time interval. Currently, the UTC and GPS times differ by 18 seconds. Such time shift can last for several hours or a full day and thus need to be corrected.

Timing errors of 1 s or smaller are not clearly evident during routine seismological analyses, but can be revealed from station “log” files, if provided by the registration system. Small time shifts can occur as a result of improper time synchronisation due to a loss of the GPS signal or due to a failure of an .oscillator tuning the station time. This third item is a more complicated issue and it allows only an approximate time reconstruction. A failure of the oscillator tuning can cause a jump or a linear increase of timing errors in data. However, such difficulties should occur only exceptionally. If they happen and we are able to identify such problems and reconstruct the real timing, it is necessary to correct times directly in the Miniseed data, which is more complex task than applying corrections in the metadata. The same concerns a reconstruction of the correct time after a loss of the time synchronisation. When checking our data, we have found an oscillator failure at station A087A, which resulted in a final time error of 0.18 s during eight days in October 2015.

Keeping exact time in seismic data is a delicate issue. But severe errors due to asynchronous application of the leap seconds or due to switches between the UTC and GPS times, can be identified and corrected automatically in any data set, including the whole AA dataset. Small time changes must be solved individually.

## 4.4 Interchange of components and polarity reversal

Sometimes results from different studies dealing with waveforms raise a suspicion that the components of seismograms are interchanged and/or the polarities reversed. Though the case is rare, we found it several times in different data sets, including data from permanent observatories. Surprisingly, the component interchange can occur during station operation, e.g., twice in the AlpArray stations until now. The simplest way to verify the correct identification of the three components is a comparison of waveforms for a selected strong teleseismic event recorded on several nearby stations, which we call the Waveform similarity method (Fig. 9a). Several other methods can be used as well, e.g., a visualization of daily means of signal amplitudes, sometimes called offsets (Fig. 9b), or a comparison of noise levels on the vertical and horizontal components in PPSD. In case of correct component identification, the noise level on the vertical component should be lower than that on the horizontal components. Correction of interchanged components can be done either in the metadata, or preferably directly in the Miniseed data.

Reversed polarity of components, arising from different technical reasons, is not as rare as one would expect. We identified polarity reversals by the manual Waveform similarity method for nearby stations. We can also use a single-station method that is based on a semi-automatic search of Rayleigh wave polarization (the Rayleigh-wave polarization-angle method) originally developed for verification of sensor orientation. Differences between the Rayleigh wave polarization and the theoretical back-azimuth are plotted againts the theoretical back-azimuths (Fig. 10). If only one horizontal component is reversed, the differences change linearly between -180° and +180°. The zero difference reflects the fact that the reversed component does not play any role in the component summation and identifies the components with the correct polarity (see Fig. 10a; the EW component is the correct one). In case of the reversed polarity on the Z component, or if both horizontal components are reversed, the differences between the Rayleigh-wave polarizations and the theoretical back-azimuths attain values around 180° for all back-azimuths (Fig. 10b). Moreover, we have also identified an interchange of both horizontal components in combination with their polarity reversals. This complicated case can be solved by combination of both methods mentioned above and by a careful analysis of the results. Similarly to the component interchange, the component reversal can be corrected either in the metadata, or preferably in the Miniseed data.

## 4.5 Gain imperfection

Anomalous signal amplitudes due to imperfectly set gains on one or more components are not very frequent in comparison with the sensor mis-orientations, but their danger for data analysis procedures is similarly large. We can recognize anomalously large or low recorded amplitudes in two ways: first, by means of technical devices, such as control and calibration units (see Sect. 3), and second, by means of software methods applied on recorded seismic signals.

One of possible software inspection of the amplitude size can be based on ambient noise, which is evaluated anyway. Moreover, ambient noise is the only continuous signal in seismic data. We have implemented a new Ambien noise gain method which compares ratios of normalized power spectra between the three components in a range of 4-8 seconds. In this range, the secondary microseisms are substantially larger than noise from local sources. Directionality of the microseisms due to different sources is eliminated by normalizing the spectrum of each trace via an average spectrum calculated over the traces of surrounding stations. The spectra are calculated within different time intervals, e.g., weeks, months, or a whole time range. Resulting ratios of the spectra provide a running record of individual channel sensitivity and allow us to follow potential changes of the amplitudes in a course of time. In combination with sporadic in-situ gain controls by the Gain and calibration box (Sect. 3.5), we have reliable control of potential anomalous size of recorded amplitudes and thus we can determine when a detected change in the gain occurred. We estimate a precision of the gain determination by the Aambient noise gain method at 1-2 dB in dependence on a length of time period analysed. Tiny variations of the curves in Fig. 11 are within this limit, but the differences between the curves are stable.

We document a successful use of the hardware and software methods on data from the two seismic experiments. During the data processing, we have found that the power spectra of the EW components at stations AAE14 (AlpArray-EASI) and A087A (AASN) are lower by approximately 11 dB (Fig. 11a). The NS/Z component ratio is close to zero, while the EW/Z and EW/NS ratios, where the EW component is involved, are 10 dB lower. Station documentations identified that stations AAE14 and A087A were equipped with an identical sensor and data acquisition system. Therefore, afterwards we tested the gain of each component of the sensor-DAS pair with the calibration boxes as described in Sect. 3. The test confirmed the amplitudes recorded on the EW component were 3.6 times smaller (20*log3.6 = 11 dB) than it should be. The technical error in the acquisition system was identified and repaired. If such error is identified by an in-situ measurement, then it can be immediately eliminated (DAS can be repaired or changed, as it was possible in case of running station A087A). Metadata of A087A for a previous period, as well as the metadata of the AAE14 station active in the finished AlpArray-EASI measurements, were corrected subsequently. In another case we have found that either the amplitudes on EW components are about two-times larger, or the gains of the NS and Z components are lower by ~6 dB at stations AAE15 (AlpArray-EASI) and A088A (AASN) (Fig. 11b). Results of the normalized PPSD ratios are only relative ones. The absolute value - the half-size gain compared with the declared one, was identified by an in-situ measurement with the use of the STS-2 control and calibration unit (see Sect. 3.4). Source of the low gain was localized in a defect cable of the seismometer. The double-checked gain levels of each component (by the hardware units and by the software calculating the normalized PPSD ratios) enabled us to correct reliably the gains in the station metadata files and thus to correct anomalous amplitudes.

## 4.6 Drift of sensor mass position

One of artefacts seen in the PPSD reflects a failure of the automatic mass re-centring of the sensor (McNamara and Buland, 2004). If a seismometer is not able to correct a drift of the mass position itself, amplitudes of seismic signals become saturated. Signal corresponding to such a time period has a characteristic “flat” spectrum shape (Fig. 12a). The flat course in an interval of ~0.3-50 s differs clearly from the shape of the noise distribution modulated by secondary microseisms. The large undesirable drift of the mass position from its central position limits the dynamic range of the sensor and therefore, it needs to be identified as soon as possible. Running information about a sensor mass “drift” comes from the MAX/MIN amplitude extremes reported by GAIA DAS in daily SMS reports (see also Fig. 13).. Besides this metric check, daily means of recorded amplitudes (Fig. 12b) serve as an independent fast and easy tool for ex-post identification of the mass centring problem. Moreover, complementing the daily amplitude means by their standard deviations and absolute values of daily amplitude extremes (maxima or minima) we can better assess the state of health of each station (Fig. 12c).

To summarize application of different methods of seismometer-GAIA DAS pair operation and recorded data quality checking, either by software or hardware tools presented above, we plot an optimal workflow in Figure 13. The scheme comes from our experience with data from several previous passive experiments. Some methods give indications about an error, which requires to be verified later. Some of the methods are repeated in time, in attempts to detect changes which can occur during station operation and thus could not be revealed by the Huddle pre-installation test.

# 5 Conclusions

We have developed both the hardware and software tools to contribute reliable high-quality waveform data to passive seismic experiments. At present, twenty broad-band stations of the Czech MOBNET pool of temporary stations are incorporated in the AlpArray Seismic Network. The stations were also deployed in the previous AlpArray-EASI complementary experiment. To assure high-degree reliability of the STS-2/CMG seismometer-DAS pairs performance, we have developed four special control devices for seismometers of different types and one for the GAIA DAS. The devices calibrate both the sensors and data acquisition systems in-situ and allow us to check the gain and the polarity of all three components. We emphasise the importance of precise sensor orientation by a gyrocompass both during station installations and of its regular checks during the field measurements. Information extracted from probabilistic power spectral density, spectra ratios, averages of daily amplitudes and of other parameters, followed by the designed procedures in routine data processing, allow us to identify several other problems, e.g., imperfectly set gains, interchange of components and polarity reversals, insufficient sensor mass centring and last, but not the least, time issues. The hardware control in-situ and the ex-post software data checking represent the double check of data quality. The former removes problems immediately in field, the latter allows restoring data back in time, until the moment when a problem occurred. The fully-automated software methods could be used for whole the AlpArray data set. We believe that the newly developed control and calibration units for setting sensor-DAS systems and the documentation of the significance of careful data-quality control with the use of the software tools, could be helpful for other groups participating in collaborative passive seismic experiments.

# Data availability

# Data from the MOBNET pool as a part of the AlpArray project is stored in EIDA (www.orfeus-eu.org/eida/), currently with restricted access (http://www.alparray.ethz.ch/research/complementary-experiments/easi/data-acess-citation/ and http://www.alparray.ethz.ch/seismic_network/backbone/data-access/).

### Team list

# György HETÉNYI, Rafael ABREU, Ivo ALLEGRETTI, Maria-Theresia APOLONER, Coralie AUBERT, Maxime BES DE BERC, Götz BOKELMANN, Didier BRUNEL, Marco CAPELLO, Martina ČARMAN, Adriano CAVALIERE, Jérôme CHÈZE, Claudio CHIARABBA, John CLINTON, Glenn COUGOULAT, Wayne CRAWFORD, Luigia CRISTIANO, Tibor CZIFRA, Ezio D’ALEMA, Stefania DANESI, Romuald DANIEL, Iva DASOVIĆ, Anne DESCHAMPS, Jean-Xavier DESSA, Cécile DOUBRE, Sven EGDORF, ETHZ-SED Electronics Lab, Tomislav FIKET, Kasper FISCHER, Wolfgang FRIEDERICH, Florian FUCHS, Sigward FUNKE, Domenico GIARDINI, Aladino GOVONI, Zoltán GRÁCZER, Gidera GRÖSCHL, Stefan HEIMERS, Ben HEIT, Davorka HERAK, Marijan HERAK, Johann HUBER, Dejan JARIĆ, Petr JEDLIČKA, Yan JIA, Hélène JUND, Edi KISSLING, Stefan KLINGEN, Bernhard KLOTZ, Petr KOLÍNSKÝ, Michael KORN, Josef KOTEK, Lothar KÜHNE, Krešo KUK, Jürgen LOOS, Deny MALENGROS,  Lucia MARGHERITI, Christophe MARON, Xavier MARTIN, Marco MASSA, Francesco MAZZARINI, Thomas MEIER, Laurent MÉTRAL, Irene MOLINARI, Milena MORETTI, Helena MUNZAROVÁ, Anna NARDI, Jurij PAHOR, Anne PAUL, Catherine PÉQUEGNAT, Damiano PESARESI, Davide PICCININI, Claudia PIROMALLO, Thomas PLENEFISCH, Jaroslava PLOMEROVÁ, Silvia PONDRELLI, Snježan PREVOLNIK, Roman RACINE, Marc RÉGNIER, Miriam REISS, Joachim RITTER, Georg RÜMPKER, Simone SALIMBENI, Detlef SCHULTE-KORTNACK, Werner SCHERER, Sven SCHIPPKUS, Vesna ŠIPKA, Daniele SPALLAROSSA, Kathrin SPIEKER, Josip STIPČEVIĆ, Angelo STROLLO, Bálint SÜLE, Gyöngyvér SZANYI, Eszter SZŰCS, Christine THOMAS, Frederik TILMANN, Stefan UEDING, Massimiliano VALLOCCHIA, Luděk VECSEY, René VOIGT, Joachim WASSERMANN, Zoltán WÉBER, Christian WEIDLE, Viktor WESZTERGOM, Gauthier WEYLAND, Stefan WIEMER, David WOLYNIEC, Thomas ZIEKE, Mladen ŽIVČIĆ

# Acknowledgements

Cooperation with participants of the AlpArray projects is greatly appreciated as well as suggestions and recommendations of both anonymous reviewers, which improved substantially the manuscript. Research of the Czech team was supported by grants No. M100121201 of the Czech Academy of Sciences and by the project CzechGeo/EPOS-Sci (CZ.02.1.01/0.0/0.0/16_013/0001800) financed from Operational Programme Research, Development and Education. Data acquisition from permanent observatories and developing the devices was supported by the project of Large research infrastructure CzechGeo/EPOS, grants Nos. LM2010008 and LM2015079. Developing the software for data-quality check was supported by grant No. COST LD15029 of the Ministry of Education, Youth and Sports. Several figures and calculations have been prepared with the use of the Generic Mapping Tools (Wessel and Smith, 1998) and ObsPy (Krischer et al., 2015).

# References

AlpArray Seismic Network: Eastern Alpine Seismic Investigation (EASI) - AlpArray Complementary Experiment, AlpArray Working Group, Other/Seismic Network, doi:10.12686/alparray/xt_2014, 2014.

AlpArray Seismic Network: AlpArray Seismic Network (AASN) temporary component, AlpArray Working Group, Other/Seismic Network, doi:10.12686/alparray/z3_2015, 2015.

Custódio, S., Dias, N. A., Caldeira, B., Carrilho, F., Carvalho, S., Corela, C., Díaz, J., Narciso, J., Madureira, G., Matias, L, Haberland, Ch., and WILAS Team: Ambient noise recorded by a dense broadband seismic deployment in Western Iberia, B. Seismol. Soc. Am., 104, 2985–3007, doi:10.1785/0120140079, 2014.

Díaz, J., Villaseñor, A., Morales, J., Pazos, A., Cordoba, D., Pulgar, J., García-Lobón, J. L., Harnafi, M., Carbonell, R., Gallart, J., and TopoIberia Seismic Working Group: Background noise characteristics at the IberArray broadband seismic network, B. Seismol. Soc. Am., 12, 618–628, doi:10.1785/0120090085, 2010.

Ekström, G., and Busby, R. W.: Measurements of seismometer orientation at USArray transportable array and backbone stations, Seismol. Res. Letters, 79, 555–561, doi:10.1785/gssrl.79.4.554, 2008.

Fuchs, F., Kolínský, P., Gröschl, G., Bokelmann, G., and the AlpArray Working Group: AlpArray in Austria and Slovakia: technical realization, site description and noise characterization, Adv. Geosci., 43, 1–13, doi:10.5194/adgeo-43-1-2016, 2016.

Handy, M. R., Schmid, S. M., Bousquet, R., Kissling, E., and Bernoulli, D.: Reconciling plate-tectonic reconstructions of Alpine Tethys with the geological–geophysical record of spreading and subduction in the Alps, Earth-Science Reviews, 102, 121–158, doi: 10.1016/j.earscirev.2010.06.002, 2010.

Kissling, E., Schmid, S. M., Lippitsch, R., Ansorge, J., and Fügenschuh, B.: Lithosphere structure and tectonic evolution of the Alpine arc: new evidence from high-resolution teleseismic tomography, in: D. Gee, R.A. Stephenson (Eds.), European Lithosphere Dynamics, Geological Society London, Memoirs 32, 129–145, 2006.

Krischer, L., Megies, T., Barsch, R., Beyreuther, M., Lecocq, T., Caudron, C., and Wassermann, J.: ObsPy: a bridge for seismology into the scientific Python ecosystem, Computational Science & Discovery, 8, 014003, doi:10.1088/1749-4699/8/1/014003, 2015.

Lippitsch, R., Kissling, E., and Ansorge, J.: Upper mantle structure beneath the Alpine orogen from high-resolution teleseismic tomography, J. Geophys. Res., 108 (B8), 2376, doi:10.1029/2002JB002016, 2003.

Molinari, I., Clinton, J., Kissling, E., Hetényi, G., Giardini, D., Stipčević, J., Dasović, I., Herak, M., Šipka, V., Wéber, Z., Gráczer, Z., Solarino, S., the Swiss-AlpArray Field Team, and the AlpArray Working Group: Swiss-AlpArray temporary broadband seismic stations deployment and noise characterization, Adv. Geosci., 43, 15–29, doi:10.5194/adgeo-43-15-2016, 2016.

McNamara, D. E., and Buland, R. P.: Ambient noise levels in the continental United States, B. Seismol. Soc. Am., 94, 1517–1527, doi:10.1785/012003001, 2004.

Peterson, J.: Observations and modeling of seismic background noise, USGS Open-File report, 93–322, 1993.

Plomerová, J., Achauer. U., Babuška, V., Vecsey, L., and BOHEMA working group: Upper mantle beneath the Eger Rift (Central Europe): plume or asthenosphere upwelling?, Geophys. J. Int., 169, 675–682; doi:10.1111/j.1365-246X.2007.03361.x, 2007.

Rueda, J., and Mezcua, J.: Orientation analysis of the Spanish broadband national network using Rayleigh-wave polarization, Seismol. Res. Letters, 86 (3), 929-940, doi: 10.1785/0220140149, 2015.

Stachnik, J.C, Sheehan, A. F., Zietlow, D. W., Yang, Z., Collins, J., and Ferris, A.: Determination of New Zealand Ocean Bottom Seismometer Orientation via Rayleigh-Wave Polarization, Seismol. Res. Letters, 83 (4), 704–713, doi:10.1785/0220110128, 2012.

Vecsey, L., Plomerová, J., Babuška, V., and PASSEQ Working Group: Mantle lithosphere transition from the East European Craton to the Variscan Bohemian Massif imaged by shear-wave splitting, Solid Earth, 5, 779–792, doi:10.5194/se-5-779-2014, 2014.

Wang, X., Chen, Q. F., Li, J., and Wei, S. J.: Seismic Sensor Misorientation Measurement Using P-Wave Particle Motion: An Application to the NECsaids Array, Seismol. Res. Letters, 87 (4), 901–911, doi:10.1785/0220160005, 2016.

Wessel, P., and Smith, W. H. F.: New, improved version of the Generic Mapping Tools released, Eos T. AGU, Volume 79, 579, 1998.

Wolin, E., van der Lee, S., Bollmann, T. A., Wiens, D. A., Revenaugh, J., Darbyshire, F. A., Frederiksen, A. W., Stein, S., and Wysession, M. E.: Seasonal and diurnal variations in long-period noise at SPREE stations: the Influence of soil characteristics on shallow stations performance, B. Seismol. Soc. Am., 105, 2433–2452, doi:10.1785/0120150046, 2015.

| Network | Station | Latitude | Longitude | Elevation (m) | Site name | Housing type | Sensor ground | Sensor | Datalogger | Start time | End time |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| XT | AAE01 | 50.6075 | 13.4320 | 590 | Hora Sv. Kateriny | adit | concrete on bedrock | STS-2 120s | GAIA-2T | 2014-07-23 | 2014-11-21 |
| . | . | . | . | . | . | . | . | STS-2 120s | Quan330S | 2014-11-21 | 2015-10-22 |
| XT | AAE02 | 50.5107 | 13.2526 | 843 | Hora Sv. Sebastiana | building/cellar | concrete | STS-2 120s | GAIA-1 | 2014-06-24 | 2015-09-03 |
| XT | AAE03 | 50.4306 | 13.4300 | 305 | Drouzkovice | building | concrete | CMG-40T 30s | GAIA-1 | 2014-06-24 | 2015-09-03 |
| XT | AAE04 | 50.3545 | 13.2588 | 388 | Uhostany | church | stone floor | STS-2 120s | GAIA-3 | 2014-07-26 | 2015-09-12 |
| XT | AAE05 | 50.2522 | 13.3696 | 301 | Krasny Dvur | castle | stone floor | STS-2 120s | GAIA-1,2T | 2014-07-01 | 2015-08-25 |
| XT | AAE06 | 50.1747 | 13.2520 | 545 | Valec | castle/cellar | stones in cement | CMG-3T 120s | GAIA-1,2T | 2014-07-01 | 2015-08-25 |
| XT | AAE07 | 50.0733 | 13.4219 | 455 | Ostrovec | church | stone floor | STS-2 120s | GAIA-1 | 2014-07-01 | 2015-09-03 |
| XT | AAE08 | 49.9910 | 13.2322 | 409 | Manetin | castle | concrete | STS-2 120s | GAIA-1 | 2014-07-17 | 2015-09-23 |
| XT | AAE09 | 49.8890 | 13.4135 | 493 | Obora | church | tiles on concrete | STS-2 120s | GAIA-1 | 2014-07-04 | 2015-10-02 |
| XT | AAE10 | 49.7998 | 13.2509 | 353 | Ceminy | castle | concrete | STS-2 120s | GAIA-1,2 | 2014-07-04 | 2015-09-12 |
| XT | AAE11 | 49.7030 | 13.4692 | 345 | Stary Plzenec | church | stone floor | STS-2 120s | GAIA-1 | 2014-06-20 | 2015-10-02 |
| XT | AAE12 | 49.6045 | 13.2629 | 360 | Dnesice | building/cellar | concrete | STS-2 120s | GAIA-1 | 2014-07-10 | 2015-10-02 |
| XT | AAE13 | 49.5289 | 13.4547 | 480 | Lazne Letiny | building | tiles on concrete | STS-2 120s | GAIA-1 | 2014-06-20 | 2015-10-02 |
| XT | AAE14 | 49.4427 | 13.2495 | 386 | Dolany | church | tiles on concrete | STS-2 120s | GAIA-2 | 2014-07-04 | 2015-09-12 |
| XT | AAE15 | 49.3648 | 13.4141 | 680 | Zdeborice | church | stone floor | STS-2 120s | GAIA-1 | 2014-07-17 | 2015-09-13 |
| XT | AAE16 | 49.2642 | 13.2193 | 643 | Depoltice | church | tiles on concrete | STS-2 120s | GAIA-1 | 2014-07-10 | 2015-09-13 |
| XT | AAE17 | 49.1554 | 13.4379 | 890 | Dobra Voda | church | brick floor | CMG-3ESP 30s | GAIA-1 | 2014-06-17 | 2015-09-14 |
| XT | AAE18 | 49.0982 | 13.2165 | 685 | Schwellhausl | cellar | concrete | STS-2 120s | GAIA-1 | 2014-08-28 | 2015-09-10 |
| XT | AAE19 | 48.9712 | 13.4825 | 1175 | Breznik | building | tiles | CMG-3ESP 30s | GAIA-1 | 2014-06-16 | 2015-08-29 |
| XT | AAE20 | 48.8896 | 13.2981 | 615 | Eppenschlag | building | concrete | STS-2 120s | GAIA-1 | 2014-08-21 | 2015-09-10 |
| Z3 | A071A | 49.7419 | 12.6911 | 502 | Stare Sedliste | church | stone floor | CMG-3T 120s | GAIA-1 | 2015-10-27 | operating |
| Z3 | A072A | 49.4683 | 13.1735 | 495 | Chudenice | castle/cellar | stones in cement | STS-2 120s | GAIA-2T,1 | 2015-08-27 | operating |
| Z3 | A073A | 49.9916 | 13.2331 | 407 | Manetin | castle/cellar | stones in cement | STS-2 120s | GAIA-1 | 2015-09-23 | operating |
| Z3 | A074A | 49.6715 | 13.5309 | 385 | Kozel | building | tiles on concrete | STS-2 120s | GAIA-1 | 2015-10-02 | operating |
| Z3 | A075A | 50.0377 | 13.8737 | 285 | Krivoklat | building | concrete | STS-2 120s | GAIA-1 | 2015-10-06 | operating |
| Z3 | A076A | 49.6168 | 14.1494 | 532 | Makova Hora | church | concrete on bedrock | CMG-3T 120s | GAIA-1 | 2015-09-08 | operating |
| Z3 | A077A | 49.2705 | 14.0739 | 370 | Kestrany | castle/cellar | protrusion wall | CMG-3ESP 30s | GAIA-1,3 | 2015-11-03 | operating |
| Z3 | A078A | 48.8640 | 14.2845 | 1060 | Klet | urban free field | concrete | CMG-3ESP 30s | GAIA-1 | 2015-10-20 | operating |
| Z3 | A079A | 49.2288 | 14.7074 | 438 | Drachov | church | stone floor | STS-2 120s | GAIA-1 | 2015-10-15 | operating |
| Z3 | A080A | 49.6840 | 14.9288 | 502 | Loreta | building | concrete | STS-2 120s | GAIA-1 | 2015-10-12 | operating |
| Z3 | A081A | 50.0752 | 15.0341 | 228 | Dobrichov | building | stones in cement | STS-2 120s | GAIA-1 | 2015-09-25 | operating |
| Z3 | A082A | 50.0610 | 15.6502 | 220 | Zivanice | church | concrete | STS-2 120s | GAIA-1 | 2015-10-08 | operating |
| Z3 | A083A | 49.6959 | 15.6077 | 573 | Cachotin | church | tiles on concrete | STS-2 120s | GAIA-1 | 2015-10-16 | operating |
| Z3 | A084A | 48.9434 | 15.7007 | 403 | Bitov | castle | tiles on concrete | STS-2 120s | GAIA-1 | 2015-09-22 | operating |
| Z3 | A085A | 49.4392 | 16.1962 | 458 | Strazek | church | tiles on concrete | STS-2 120s | GAIA-1 | 2015-10-07 | operating |
| Z3 | A086A | 49.8528 | 16.1457 | 391 | Nove Hrady | castle/cellar | stone on bricks | STS-2 120s | GAIA-1 | 2015-09-19 | operating |
| Z3 | A087A | 49.7049 | 16.8893 | 430 | Bouzov | castle | tiles on concrete | STS-2 120s | GAIA-3 | 2015-09-19 | operating |
| Z3 | A088A | 49.4303 | 17.2911 | 211 | Tovacov | castle | concrete | STS-2 120s | GAIA-1 | 2015-09-24 | 2017-04-19 |
| . | A088B | 49.4305 | 17.2919 | 208 | . | . | stones in cement | . | . | 2017-04-19 | operating |
| Z3 | A089A | 49.1521 | 17.0920 | 263 | Nesovice | castle/cellar | concrete | STS-2 120s | GAIA-1 | 2015-11-08 | operating |
| Z3 | A090A | 49.3655 | 17.8278 | 659 | Maruska | underground shelter | concrete in soil | CMG-3ESP 30s | GAIA-1 | 2015-09-24 | 2015-12-02 |
| . | . | . | . | . | . | . | . | CMG-40T 30s | GAIA-1 | 2015-12-02 | 2016-04-06 |
| . | . | . | . | . | . | . | . | CMG-3ESP 30s | GAIA-1 | 2016-04-06 | operating |

Table 1: List of Czech temporary stations from the MOBNET pool involved in the AlpArray-EASI complementary experiment (network code XT) and the AlpArray Seismic Network (AASN, network code Z3).

| EASI stations (CZ) | 1st measurement [deg] | 2nd measurement [deg] | 3rd measurement [deg] | Difference [deg] |
| --- | --- | --- | --- | --- |
| AAE01 | 359.9 | 359.9 | x | x |
| AAE02 | 7.4 | 0.4 | 0.4 | 0 |
| AAE03 | 4.0 | 359.8 | 0.7 | 0.9 |
| AAE04 | - | - | 340.8 | - |
| AAE05 | 357.1 | 359.3 | 0.6 | 1.3 |
| AAE06 | 3.2 | 0.9 | 0.9 | 0 |
| AAE07 | 355.6 | 0.4 | 2.6 | 2.2 |
| AAE08 | 358.0 | 0.6 | 0.8 | 0.2 |
| AAE09 | 2.1 | 0.3 | 359.6 | -0.7 |
| AAE10 | 8.7 | 0.4 | 3.1 | 2.7 |
| AAE11 | 5.2 | 0.7 | 359.5 | -1.2 |
| AAE12 | 2.9 | 0.7 | 359.5 | -1.2 |
| AAE13 | 282.0 | 0.7 | 352.3 | -8.4 |
| AAE14 | 2.3 | 359.5 | 357.2 | -2.3 |
| AAE15 | 3.2 | 359.9 | 359.4 | -0.5 |
| AAE16 | 2.2 | 359.4 | 0.5 | 1.1 |
| AAE17 | 6.2 | 0.2 | 0.8 | 0.6 |
| AAE18 | 7.2 | 0.4 | 6.9 | 6.5 |
| AAE19 | 6.0 | 359.8 | 359.6 | -0.2 |
| AAE20 | 3.5 | 0.2 | 0.4 | 0.2 |

Table 2: Sensor orientation during the AlpArray-EASI complementary experiment measured by Gyrocompass. The largest differences are in bold.

1st measurements: checking original towards North orientation of seismometers determined during station installation with the use of a standard compass

2nd measurements: towards the North reorientation of sensors with the gyrocompass

3rd measurements: check of sensor orientation at the end of registration

Difference:  difference between the 3rd and 2nd measurements

Figure 1: Broad-band stations of the MOBNET pool involved in the AlpArray project: (a) in the network of the AlpArray-EASI complementary project (2014-2015) and (b) in the AlpArray seismic network (AASN). The broad-band stations of the complete AASN cover the area within 250 km distance (outer white curve) from the smoothed 800 m altitude line of the Alps (inner white curve). The array consists of 623 stations including the OBS in the Ligurian Sea.

Figure 2: Schematic map of major tectonic units of the Bohemian Massif and seismic stations of the MOBNET pool involved in the network of the AlpArray-EASI complementary project (blue circles) and the AlpArray Seismic Network (AASN, green circles) of the Alp[Array project. CRSN – Czech Regional Seismic Network, WEBNET - West Bohemian Network.

Figure 3: Example of installation of one of the broad-band MOBNET stations in the AlpArray-EASI complementary experiment. Besides the photo-documentation of the installation of the AAE08 station (Manětín), we show its location (upper right), station coordinates and period of recordings (top panel), describe the installation of accessories (GPS antenna) and the bedrock (left panel in the middle). The central panel shows probabilistic power spectral density of ambient noise on the vertical component during the first five months of the station registration.

Figure 4: Example of installation of one of the broad-band MOBNET stations in the AlpArray Seismic Network (AASN).. Besides the photo-documentation of the installation of the A076A station (Maková Hora), we show its location (upper right), station coordinates, start time of recordings (top panel), describe the installation including accessories (GPS antenna) and the bedrock (left panel in the middle). The central panel shows probabilistic power spectral density of ambient noise during the first four months of the station registration.

Figure 5: Probabilistic power spectral density (PPSD) on the HHE horizontal component of the AAE01 station of the AlpArray-EASI experiment. Branching of the PPSD reflects both diurnal and seasonal variations of noise and it is clearly visible both in the high-frequency and in the low-frequency bands, respectively. Noise level during the summer period is low and far below the required noise level (20 dB below the New High Noise Model (NHNM – upper grey curve; Peterson, 1993)) even for periods larger than 10 s, whereas during winter time, noise exceeds the limit of 10 dB below NHNM for periods larger than 30 s

Figure 6: Data completeness of MOBNET stations in the AlpArray projects: (a) in the AlpArray-EASI complementary field measurements and (b) in the ongoing AlpArray experiment.

Figure 7: Control and calibration units developed for the broad-band seismometers and the GAIA DAS to guarantee the high-quality of recorded data.

Figure 8: Medians of probabilistic power spectral density (PPSD) of seismic noise at the MOBNET stations involved (a) in the AlpArray-EASI complementary experiment and (b) in the AlpArray Seismic Network (AASN). The grey thick lines mark the New Low and High Noise Models (NLNM and NHNM; Peterson, 1993). The thick red line (AAHNL) marks the upper limit of the recommended noise level below the NHNM.

Figure 9: The Waveform similarity method showing the interchange of the EW and Z components (a) at the AAE05 station. Daily amplitude means (b) of all three components at the AAE05 station during June – July 2015 indicate the interchange of all three components. Vertical dashed lines show dates of station servicing.

Figure 10: Rayleigh-wave polarization-angle method showing differences between the Rayleigh wave polarizations and theoretical back-azimuths (BAZ) in dependence on BAZ of earthquakes: (a). linear dependence of the difference identifies the polarity reversal of the NS component. Polarization of waves arriving from the East or West is not affected by the reversed NS component. Station PA10 of the PASSEQ experiment 2006-2008 (Vecsey et al., 2014) is used as an example; (b) the constant difference of ~180° corresponds to the polarity reversal of both horizontal components at the A077A station of AASN.

Figure 11: Seismometer gain check by the Ambient noise gain method. Gain imperfection can be identified from monthly averages of normalized power spectral ratios. Both panels show that either gain of the EW component or, gains of both NS and Z components are not correct. To decide which amplitudes are correct and which are not, one has to check gain of each component of the sensor-DAS pair with the calibration units, as described in Section 3.

Figure 12: Error in sensor mass centring. Flat curves in the probabilistic power spectral density (PPSD) (a) reflect an amplitude saturation due to insufficient mass centring and identify an existence of the problem. The time range of the failure of the automatic mass re-centring is determined by changes of daily means of amplitudes (b), their standard deviations (c) and absolute values of daily amplitude extremes (d)..

Figure 13: Optimal workflow of temporary station control and data-quality checks to assure archiving the high quality data. Hardware and software procedures are in rectangular and oval boxes, respectively.